## Slide 1: Connect the hardware
Notes: Now that we’ve done the calculation and simulation let’s set up the hardware and measure the performance. First set the jumpers on the PLABS hardware as shown. In this experiment, no external signal source is required as the PLABS board has an on board precision DC source. Next plug the PHI into the PLABS channel 2 connection and install the OPA320_Noise1 coupon board into the channel 2 socket. Finally plug the USB cables into your computer. When connecting the USB cable to the computer an LED on the PHI should blink. Let’s pause for a moment to set the jumpers and connect the USB cable.

## Slide 2: Measured vs Expected Results
 |  | Vin = 0.1V DC, fsamp = 1MHz, Vref =5V, LSB = 76.29uV |  |  |  |  |  |  | 
Device | Hand Calc | Simulated |  |  |  | Example Measurements |  | Your Measurements | 
 | Amp Noise (uV) | Amp Noise (uV) | ADC Noise (uV) | Ref Noise (uV) | Total Noise (uV) | Sigma | Noise (uV) | Sigma | Noise (uV)
Noise1 | 196.7 | 170.8 | 39.6 | 11.2 | 175.4 | 2.18 | 166.3 |  | 
Noise2 | 99.1 | 78.8 | 39.6 | 11.2 | 88.4 | 1.16 | 88.5 |  | 
Your results should show the same trend as the expected result but the specific values will differ.
Notes: This table summarizes the calculated, simulated, and measured results for this experiment. Your measured results should be close to the example measurement, but you may see some difference. Remember, to translate Sigma to an RMS noise voltage you multiply by the LSB bit width. Comparing the total simulated noise to the example measurement shows that the two numbers compare very well. Remember, there are many factors that impact noise, so you should not expect the results to match exactly. Pause now and fill in your measurements.