## Slide 1: Understanding and Calibrating the Offset and Gain for ADC Systems TIPL 4202 TI Precision Labs – ADCs
Created by Dale Li, Art Kay
Presented by Peggy Liska
Notes: Hello, and welcome to the TI Precision Lab covering data converter error sources. Overall, this video will discuss how gain and offset errors can be calculated and eliminated through calibration. We will start by doing an offset and gain error calculation for a data converter system. Next, we will discuss a few methods for calibrating out this error. And finally, we will introduce some different error sources that are difficult to calibrate out.

## Slide 2: Offset Voltage Calculation
Typical offset at ADC Input
Typical offset at ADC Input
Device | PARAMETER |  | MIN | TYP | MAX | UNITS
LMP8481 | VOS | Offset | -265 | ±80 | +265 | µV
OPA320 | VOS | Offset | -150 | ±40 | +150 | µV
ADS8860 | EO | Offset | -2 | ±0.8 | +2 | mV
Number of Standard Deviations | Probability Inside Limit
±1∙σ | 68.27%
±2∙σ | 95.45%
±3∙σ | 99.73%
±4∙σ | 99.9937%
±5∙σ | 99.99994%
±6∙σ | ≈100%
Notes: In this slide, we repeat the offset calculation from the Precision Labs video titled Statistics Behind Error Analysis. Here we are looking at the typical offset for each of the two amplifiers and the ADC in the signal chain. All the offset values are referred to the input of the data converter so the gain of each amplifier needs to be taken into account. In this example, U1 has a gain of 20 so its offset is multiplied by this factor. All three offsets are combined using the square root sum of the square because they are uncorrelated Gaussian distributions. The total offset here represents a typical offset for the system which is plus or minus one standard deviation. To set a worst case limit, we can multiply the standard deviation by the appropriate factor from the table according to the system requirements. For example, if we set the system maximum specification to ±3 standard deviations, then 99.73% of the population will be inside that limit and 0.27% of the population will fail the limit. Depending on the requirement, a more conservative limit can be set.

## Slide 3: Gain Error Calculation
Device | PARAMETER |  | MIN | TYP | MAX | UNITS
R1 | ER | Tolerance | -0.1 |  | +0.1 | %
LMP8481 | EG | Gain Error | -0.6 |  | +0.6 | %
ADS8860 | EG | Gain Error | -0.01 | ±0.005 | +0.01 | %
Absolute Worse Case Gain Error
Statistical Worse Case Gain Error
Absolute Worse Case Gain Error
Statistical Worse Case Gain Error
Notes: Here we show a similar statistical analysis can be done for gain error. In this case, the current shunt resistor, current shunt amplifier, buffer, and data converter will contribute gain error. We will ignore the gain error of the buffer as it is very small and related to the open loop gain of the amplifier. In this case, some of the parameters don’t have a typical value, so we have to use the worst case for the analysis. The equations at the top of the slide show an absolute worst case analysis where the errors are directly added, and a statistical worst case where the errors are added as the square root sum of the squares. The statistical worst case is a more reasonable estimate of worst case, whereas the absolute worst case is more conservative. In this example, the error for the current shunt amplifier, U1, is the dominant source of error. For many circuits, the gain is set by external discreet resistors. For this type of circuit a Monte Carlo analysis is a good approach for finding gain error. The Precision Labs video titled Gain Error and Monte Carlo Analysis covers this topic.

## Slide 4: Offset and Gain Calibration: two test signals
Two Precision inputs applied
Average code read
Make sure amplifiers are in linear range
Notes: The transfer characteristic for most signal chains is a linear function in the form y=mx +b. Technically, there will be some nonlinear terms, but assume linearity is a good first order approximation. Offset and gain calibration are based on the idea that we can solve the straight line equation for the slope and intercept. Note that the slope error is the gain error and the intercept is the offset. Applying two different input signals and measuring the associated output code will allow you to solve for the slope and intercept. However, you have to be careful to make sure that the amplifiers are all operating in the linear region of the curve. Looking at the input output relationship you can see how it would not be possible to determine the slope of the transfer function in the non-linear region. In this example we apply 0A and 20A and measure the associated output codes. Note that 0A drives the output of U1 to 0.5V and 20A drives the output to 4.5V. So these test signals keep the system in the linear range. It is very important that the test signals used are very accurate. Any error in the test signal will introduce error in the calibration coefficients and minimize the effectiveness of the calibration. After calibration, the measured slope and measured offset, called calibration coefficients, are generally stored in the microcontroller’s memory. These coefficients are then used during normal device operation to compensate for the gain and offset error. On the next slide we will look at the math for this example.

## Slide 5: Calibration Example
Example calculation for Offset and Gain Error: | 
 | Apply 0A
 | Apply 20A
 | Compute Slope based on codes
 | Offset
 | Example Input
 | Output code
 | Uncalibrated 2mV error
 | Calibration eliminates error
Example calculation for Offset and Gain Error: | 
 | Apply 0A
 | Apply 20A
 | Compute Slope based on codes
 | Offset
 | Example Input
 | Output code
 | Uncalibrated 2mV error
 | Calibration eliminates error
Notes: This is an example calibration based on the circuit from the previous slide. In this case, the calibration test signals are at 0.5V and 4.5V. Notice that the ideal graph is shown in blue and the measured graph is shown in red. The measured graph has an offset and different slope than the ideal function. The slope can be calculated by taking the change in output code divided by the change in input voltage. The offset is determined by solving the y = mx + b equation for “b” and substituting one input and its associated output signal. Once you have the offset and slope you can correct the error for any input. In this example, we apply a 2V input signal. The uncalibrated reading for this is 2.002V, so there is a 2mV error from gain and offset. Applying the correction using the calibration coefficients, we can eliminate the error and determine the actual input signal of 2V.

## Slide 6: Offset calibration with short
Notes: Some calibration schemes apply external calibration signals during initial production to calibrate the system. In other cases, the system will have on-board precision references to generate the calibration signal. However, for some applications generating the precision calibration input signals may not be practical because the cost may be too high. In the example we just considered, the calibration signals were 0A to 20A. Generating a precision 20A signal to calibrate the system is challenging and expensive.
One way to simplify a calibration scheme is to do an offset only calibration. The nice thing about the offset calibration is that it can generally be done by shorting the input to ground. Shorting the input to ground provides a very accurate 0V input signal. The 0V input signal does not have the accuracy and drift errors inherent in a typical reference circuit that would be used for a two point gain and offset calibration. The circuit here shows a simple calibration scheme where the input signal can be disconnected and the input can be grounded for offset calibration. When 0V is applied to the input, the offset is directly read to be -30 codes. This offset includes the ADC offset as well as the offset for amplifiers U1 and U2. Also, this calibration can be run periodically to account for offset drift. Ideally, it would be good if we could also calibrate the gain error, but this would require a precision calibration source. So in order to minimize cost and complexity, some systems use this simple one-point calibration. It is important to realize that this type of calibration can only be done on an ADC with a bipolar range or a differential input range. In the next slide, we will see why this method cannot be used on unipolar ADCs.

## Slide 7: Offset Error – Unipolar
Notes: This slide shows the effect of positive and negative offset on a unipolar data converter. By “unipolar” we mean that the input signal is always positive; that is, it ranges from zero to the full-scale range, denoted as FSR. The output codes for this example range from 000H to FFFH. Although there are no negative output codes for a unipolar converter, there can be a negative offset. The graph on the left shows how negative offset would affect the ADC transfer function. The ideal curve is shown in blue and the measured curve is shown in red. Notice that the measured curve is shifted downward by the negative offset, but the transfer function is truncated at 000H. So, for this example applying a 0V input would generate a 000H output code even though the actual offset error is -003H. Therefore, you are not able to calibrate for a negative offset using a 0V calibration signal for unipolar ADCs. However, as shown on the right hand curve, you can measure a positive offset using a 0V input for unipolar ADCs. In this example, applying 0V to the input you would measure an offset of +003H. Nevertheless, this doesn’t really help from a calibration perspective as the offset of an ADC will always range from negative to positive values. In the next slide we will see why a bipolar ADC offset calibration works for a 0V input.

## Slide 8: Offset Error – Bipolar, or Unipolar with Differential input range
Notes: Here we show the offset error for a bipolar ADC, or a ADC with a differential input range. The term “bipolar” means that the input can accept both negative and positive voltages. These curves can also apply to a unipolar input that has a differential input range. The example shown earlier used the ADS9110, which is a unipolar device that has a differential input range of ±Vref. The range starts at the Negative Full Scale, denoted as NFS, and ends at the Positive Full Scale, denoted as PFS. In this case, it is possible to apply zero volts to the input and directly read the offset. Shorting the inputs to determine offset is often used as a simple approach for measurement and calibration of offset.

## Slide 9: Automatic Offset Calibration: ADS7042
Inputs disconnected during calibration
Digital command initiates calibration
Offset correction stored in register and automatically subtracted after each conversion
PARAMETER | MIN | TYP | MAX | UNIT
Uncalibrated Offset |  | ±12 |  | LSB
Calibrated Offset | -3 | ±0.5 | +3 | 
Notes: Some data converters have a convenient feature that allows automatic offset calibration. In general, this calibration is initiated by sending the ADC a digital command. This can be done periodically so that offset drift is also accounted for. During calibration, the ADC’s inputs are disconnected from the rest of the circuit, so special calibration test signals are not required. Note that in this case, the calibration only applies to the data converter and offset errors from the rest of the signal chain are not corrected for in this automatic calibration. The digital value for the offset is stored in a register after calibration. The offset register is automatically subtracted from readings after calibration to correct for the offset error. This correction can significantly reduce offset error. For example, the ADS7042 typical offset is reduced from ±12 LSB to ±0.5 LSB through calibration. Finally, it is important to note that this automated calibration corrects for offset error but does not correct for gain error.

## Slide 10: Error Sources that are difficult to Calibrate
Temperature Drift
Non-linearity
Long term shift (Aging)
Hysteresis
Noise
Integral Non-linearity
Long term shift
Temperature Drift
Notes: Gain and offset errors are two common error sources that can be eliminated using calibration. Some other error sources can be difficult or impossible to calibrate out. A few examples are noted here. Offset and gain error can have a temperature drift. Integral non-linearity is a measurement of the deviation of the transfer function from an ideal transfer function. Long term shift, or aging, is a measurement of how the performance of the device degrades with time. Hysteresis indicates how the behavior of the device can be changed by temperature cycling the device. This is different from temperature drift in that the room temperature operation of the device can be altered, for example, by cycling the temperature from hot to cold and back to room temperature. After the temperature cycle, the offset and gain can be shifted by stress induced inside the device from the temperature extremes. Note that these types of errors are very difficult to correct for in calibration, so ideally they should be low compared to your system’s error requirements.

## Slide 11: Thanks for your time! Please try the quiz.
Notes: That concludes this video – thank you for watching! Please try the quiz to check your understanding of this video’s content.