## Slide 1: Calculating the total noise for ADC systems TIPL 4204 TI Precision Labs – ADCs
Created by Art Kay, Dale Li
Presented by Peggy Liska
Notes: Hello, and welcome to the TI Precision Lab covering noise calculations for ADC systems. Overall, this video will walk through how to predict the overall system noise for a data converter, amplifier, and reference. We will use SPICE and data sheet specifications to do this analysis. Finally, we will see how measured results match the predicted noise closely.

## Slide 2: SNR of Amplifier + ADC: General Equations
 | Solve for noise
 | From ADC data sheet
 | Total RMS Noise
 | ADC+Amp+Ref
 | Solve for noise
 | From ADC data sheet
 | Total RMS Noise
 | ADC+Amp+Ref
Notes: In this section we will do a calculation to find the total noise for this signal chain. In this example, the amplifier, data converter, and reference will all contribute some noise. The data converter’s noise is defined by the signal to noise ratio, but the amplifier and reference noise are normally calculated or simulated in volts rms. In order to combine the noise from all the components, we need to translate the ADC noise into volts rms. The equation on the top is the general equation for SNR. Solve this equation for ADC noise denoted as VnADC. Then add all three noise components as the square root sum of the square. Finally, it is useful to translate the total noise back to a signal to noise ratio as this will directly relate to our measured results.

## Slide 3: Find the REF6050 Noise
1. Analysis> Noise Analysis
2. Enter the bandwidth and select the diagrams
The integrated noise is the “total noise”. Look at the final value
VnRef ≈ 6.31uV rms
Notes: One way to get the total integrated rms noise for a reference is to do a simulation. To do this from TINA SPICE you first select noise analysis. Enter a frequency range that allows the integrated noise to converge to a final value. For most references, 1MHz or less should be sufficient. The total noise is the integrated rms noise at the output of the reference. In this case the integrated noise is about 6.31uV rms. The “total integrated noise” in rms is read where the Total Noise plot flattens out at high frequency to a consistent value. For this device the data sheet also provides a typical “total integrated noise” spec for a 22uF capacitor of 5uV rms, so the simulation and data sheet specification are reasonably close.

## Slide 4: Simulating Amplifier Noise
For more information: http://www.ti.com/precisionlabs
Notes: The noise for the amplifier configuration can also be simulated using TINA spice. Before running the noise analysis, it is useful to test the dc operation of the amplifier. In this case, the amplifier input is set to 0.1V so the output should be the gain times the input or 11*0.1V = 1.1V. Doing a “DC Analysis>calculate nodal voltages” confirms that the output is 1.1V and the circuit is functioning properly. A common mistake here is to ground the input. With 0V in, the amplifier’s output will try to drive to ground which is not a linear operating condition. Noise simulations require linear operation to function properly. Note that to do a Noise Analysis an AC Source must be connected to the op amp circuit. The Vin source for this circuit needs to be an AC Voltage Source with a DC Value of 0.1V.
Once linear operation of the amplifier is confirmed, run “Analysis>Noise Analysis”. Enter a frequency range that will allow the noise to converge to a final value. Normally this is about a decade beyond the bandwidth of the amplifier. In this example the amplifier’s bandwidth is 20MHz so we try 100MHz for the maximum frequency. You can see over at the right the spectral density and integrated noise curves. Notice that the integrated noise converges at 128.78uV. This is the total rms noise that we will use in the final noise calculation. If you are not familiar with noise simulation, you can learn more about this in the TI Precision Labs Op-Amps videos. A link is given at the bottom of the page for your reference.

## Slide 5: SNR of Amplifier + ADC: Example Calculation
V
Notes: Here we do the total noise calculation for the signal chain. First we need to translate the ADC SNR specification to an rms noise voltage. The equation given for the ADC noise require the full scale rms input voltage. The full scale range is 5V for this converter. The full scale range corresponds to the peak-to-peak input. To convert this to rms we divide by two and multiply by 0.707. For this example the rms input is calculate as 0.5*5V*.707 = 1.767V. The data sheet typical SNR for this device is 93dB. Applying all these factors to the noise equation yields a total noise of 39.6uV rms for this device. Next we combine the ADC, amplifier, and reference noise using the square root sum of the squares. In this example the total noise is 134uV rms. Finally, if we use the calculated total noise in the SNR equation we can calculate a modified SNR for the example that includes the amplifier and reference noise. In this example the modified SNR is 82.4dB. Thus, the ADC SNR was degraded from 93dB to 82.4dB by the amplifier and reference noise. Note that the reference noise was negligible in this case.

## Slide 6: Signal Chain Noise: Analog Engineer’s Calculator
1. Enter the total rms noise from the signal chain.
2. Enter the noise from the ADC data sheet.
3. Press “ok” and the total noise is displayed here.
http://www.ti.com/tool/analog-engineer-calc
Notes: The “Analog Engineer’s Calculator” introduced in other TI Precision Labs videos has a tool that will automatically do the calculation from the previous slide. First enter the amplifier rms noise, second enter the ADC data sheet information, and finally, press “ok” and the results are displayed to the right. Notice that these numbers match the calculation in the previous slide. The URL where you can download this calculator is at the bottom of the page.

## Slide 7: SNR of Amplifier + ADC: Measured Result
from the previous slide
rms noise is one sigma
Notes: Here we measure the noise for the example from the previous slide. One way to measure noise is to look at the histogram for a precision dc input signal. Ideally, this histogram should have one code bin that corresponds to the input voltage. In reality, however, most converters will have a Gaussian distribution of output codes that correspond to the output noise. The standard deviation of this distribution is equal to the total rms noise. In this measurement example the standard deviation, or sigma, for the distribution is 1.78 codes. We can convert this to a noise voltage by multiplying the standard deviation by the LSB width. In this example the noise is calculated to be 1.78*76.29uV which is equal to 136uV rms. The measured noise of 136uV compares very well with the calculated 134uV from the previous slide.

## Slide 8: Averaging to Reduce Noise
 | Measured RMS codes | Calculated RMS codes
Standard Deviation Raw Data | 1.80 | na
Standard Deviation 10 x Averaging | 0.59 | 0.57
Standard Deviation 100 x Averaging | 0.18 | 0.18
Measured vs. Calculated Averaging
Notes: One way to reduce noise is to optimize the signal chain by choosing low noise amplifiers and carefully selecting components to minimize noise. However, once the signal chain is optimized your may want to further reduce the noise. Averaging is another approach for reducing total noise. This example shows real world measured noise from the previous example. For a random uncorrelated Gaussian noise distribution, the total noise after averaging is the noise before averaging divided by the square root of the number of averages. The example calculation here shows the effect of averaging by a factor of 10 on the previous example. The averaged noise is by taking the 1.8 codes of un-averaged noise and dividing by the square root of 10, which yields 0.57 codes of rms noise after averaging. A 10 point rolling average was done on the ADC data, and the standard deviation of the averaged data was calculated to be 0.59. So the effect of averaging on the actual measured data very closely matched the predicted noise by dividing by the square root of N. Then, a similar experiment was done with a 100 point rolling average and both the calculated and measured averaging corresponded quite well reducing the noise to about 0.18 codes rms. Be careful as continuing to average the noise will not necessarily result in noise reduction. Averaging assumes a Gaussian uncorrelated distribution and some forms of extrinsic noise will not meet this criteria. Furthermore, there are practical limits on how low the noise can go; it cannot be averaged below one LSB bit width.

## Slide 9: Thanks for your time! Please try the quiz.
Notes: That concludes this video – thank you for watching! Please try the quiz to check your understanding of this video’s content.