## Slide 1: Voltage Reference Introduction TIPL 4501 TI Precision Labs – ADCs
Created by Luis Chioye, Art Kay
Presented by Cynthia Sosa
Notes: Hello, and welcome to the TI Precision Lab series covering voltage reference considerations for SAR ADCs. The goal of this section is to cover reference specifications, gain a deeper understanding of the SAR voltage reference behavior, and develop methods for driving the reference input that minimize error.
This video starts with an introduction on voltage references. In this video we cover a brief overview of voltage reference specifications including initial accuracy, long term stability, temperature drift, and noise. In follow on videos we will cover topics related to driving the SAR reference input. Before we begin looking at reference specifications, lets cover some more detail on the objective of the entire reference series.
.

## Slide 2
Agenda for the voltage reference series
Reference Performance Specifications:
  Initial Accuracy, Drift, Long Term Drift, and Noise
SAR REF Drive Circuit Topologies:
Standalone Reference vs. Buffered Reference
  SAR ADCs with Internal Reference Buffer
SAR REF Input Overview: The Capacitive DAC (CDAC)
Build TINA REF Input Model for a SAR:
Discrete Charge Model
TI Device Specific Model
SAR REF Drive Circuit Design:
Reference Bypass Capacitor
Reference Buffer Stability and Compensation
Notes: This slide shows the agenda for the entire series on the SAR ADC reference. In this video we will cover the first topic: initial accuracy, drift, long term drift, and noise. Topics for the other videos are given below. Without further delay, let’s get started on reference specifications.

## Slide 3: Series or Shunt Reference
Shunt
Series
Used in ADC reference input drive
Output amplifier can source and sink current
Low dropout
No external resistor required.
Cannot sink current
Limited ability to respond to load transients
Used for wide Vin and Vin transients
Can be used as floating reference.
Notes: Before starting to review different reference specifications we should cover the main differences between the two different types of voltage references: the shunt and series reference. The shunt reference acts very similar to a Zener diode. Current through the shunt reference is set by a series resistance. The shunt reference does not have an amplifier buffered so it cannot sink current. The amount of current it can source is related to the shunt resistor. Because of their limited ability to respond to load transients, these references are generally not used when driving switched capacitor reference inputs. One possible exception would be cases where the shunt reference is followed by a wide bandwidth buffer.
The series reference, on the other hand, is the type that we will be focusing on in this presentation because they are the most common type used for ADC input drive. The series regulator has an internal bandgap reference and an amplifier that buffers the reference voltage. Normally, the series regulator output can source or sink current which makes this type of regulator good for responding to transient current demands. Later in this seminar we will learn about characteristics that make some series references better than others for driving references. For now suffice it to say that we will be focusing on series voltage reference the remainder of this presentation.

## Slide 4: Initial Accuracy
At TA = 25⁰C, ILOAD = 0, CL = 1µF, and VIN = (VOUT + 0.2V) TO 18V, unless otherwise noted. |  |  |  |  |  | 
PARAMETER |  | TEST CONDITIONS | MIN | TYP | MIN | UNIT
VOUT | Output Voltage | REF5020 |  | 2.048 |  | V
 |  | REF5025 |  | 2.500 |  | 
 |  | REF5030 |  | 3.000 |  | 
 |  | REF5040 |  | 4.096 |  | 
 |  | REF5050 |  | 5.000 |  | 
 |  | REF5010 |  | 10.00 |  | 
 | Initial Accuracy: High Grade |  | -0.05 |  | 0.05 | %
 | Initial Accuracy: Standard Grade |  | -0.1 |  | 0.1 | %
Valid at the test conditions: see top of table.
Does not include effects from temperature drift, soldering, long term shift, or thermal hysteresis.
Notes: Initial accuracy is the error in the specified output voltage of the reference.
Click The error is only valid under the specific set of test conditions that the initial error is defined for.
Click
In this example, the initial accuracy is defined by the test conditions at the top of the page. That is, TA = 25⁰C, ILOAD = 0, CL = 1µF, and VIN = (VOUT + 0.2V) TO 18V, unless otherwise noted.
Notice that some references have multiple grade outs.
Click In this example, the high grade has 0.05% error and the low grade has 0.1% accuracy.
Click Note that the initial accuracy does not includes the effects of accuracy shifts from soldering, long term shift, or hysteresis. We will discuss these error sources in a moment.

## Slide 5: Initial Accuracy and Solder Shift
The solder shift is for a specific test condition:
  Example: PCB thickness, area, and soldering profile.
  Other conditions (e.g. multiple reflows) could have larger shifts.
Notes: This slide shows the histogram for initial accuracy as well as the error introduced by solder shift. The solder-heat shift distribution error is caused by stress introduced in the die during the soldering process. This error is in addition to the initial accuracy error. This distribution was measured under very specific test conditions such as the PCB thickness, area, and soldering profile. Changing the soldering conditions can impact the solder shift. For example, many PCB designs will require more than one reflow for top and bottom sides of the PCB. This can make the solder-heat shift even larger.

## Slide 6: Long Term Shift (stability)
At TA = 25⁰C, ILOAD = 0, CL = 1µF, and VIN = (VOUT + 0.2V) TO 18V, unless otherwise noted. |  |  |  |  | 
PARAMETER | TEST CONDITIONS | MIN | TYP | MIN | UNIT
LONG-TERM STABILITY: VSSOP-8 | 0 to 1000 hours at 25⁰C |  | 125 |  | ppm
LONG-TERM STABILITY: VSSOP-8 | 1000 to 2000 hours at 25⁰C |  | 45 |  | ppm
LONG-TERM STABILITY: MSOP-8 | 0 to 1000 hours at 25⁰C |  | 100 |  | ppm
LONG-TERM STABILITY: MSOP-8 | 1000 to 2000 hours at 25⁰C |  | 50 |  | ppm
96 units tested.
Caused by curing of the molding compounds, metal migration, and diffusion.
The initial shift (first few months) makes up the majority of change. Due to Self curing of molding compound.
Notes: Long term shift, sometimes called long term stability, is a measurement of how the accuracy of the devices changes with time. This is sometimes called the aging of the product. All devices will have long term shifts from ageing. However, this error is especially important in references.
The table shows the typical shift of the reference output for two different packages: VSSOP-8 and MSOP-8. The table specifies the output drift from 0 to 1000hours and from 1000 hours to 2000 hours.
Click
In this example, 96 units were tested.
click
Long term shifts are caused by curing of the molding compounds, metal migration, and diffusion. In the life of a product the long term shifts are largest initially and begin to diminish over time. The reason for the large initial shift is the self curing of the molding compound in the device package. One way to minimize the effects of long term shift is to periodically calibrate the product. Thus, the long term shifts are accounted for. Another approach is to do an initial “burn-in”. This means that the product is baked at high temperature to accelerate the ageing of the product and get past the “initial shift”.

## Slide 7: Temperature Drift (Box Method)
Box Method | 
Green | 
Red | 
Box Method | 
Green | 
Red | 
Change in output voltage vs temp.
Box method: drift over the entire range.
Instantaneous drift can be higher or lower.
If example limit is 20ppm/⁰C, red fails and green passes.
Notes: Temperature drift is the change in output voltage over temperature. This drift is normally very nonlinear for references as it uses design techniques and trim to get a relatively flat temperature response, so the residual error is nonlinear. The box method is used to calculate temperature drift. This method looks for the largest change in output voltage over temperature. The drift in any particular region can be smaller or larger than the specification. In the example, you can see that the drift from 0C to 50C is relatively flat where as the drift at low and high temperature is much higher. The box spec gives the drift across the entire range. Also note that the drift is only tested at a few key points. In this example, it is tested at -50C, 25C, and 125C. It is possible that a minima or maxima could occur at temperatures other than the test points, but the test points were selected according to the expected curvature of the drift to avoid this. In this example the limit is set to 20ppm/C, so the green curve passes and the red curve fails the limits. Temperature drift is considered a very important specification because it is usually not practical to calibrate out this error. Typically, to achieve the best temperature drift you simply choose the product with the best drift specification.

## Slide 8: 0.1Hz to 10Hz Noise: 1/f noise (flicker noise)
Measured using a special band pass filter and oscilloscope.
Not practical to filter 1/f.
Looks like temperature drift.
Good relative comparison of 1/f noise for different references.
Notes: 0.1Hz to 10Hz noise is an industry standard way of specifying 1/f or flicker noise. It is measured using an active band pass filter that attenuates signals below 0.1Hz and above 10Hz. It is not practical to filter 1/f noise as it extends down to zero frequency and you cannot set a filter cutoff to zero hertz. Because this noise is very low in frequency it can be difficult to distinguish from ambient temperature drift. The only practical way to minimize 1/f noise in references is to simply choose a reference model with good specifications. Because this error source cannot be calibrated out it is often a key consideration when selecting devices.

## Slide 9: Broadband Noise
1/f noise
Broadband noise minimized with filtering
Broadband noise
Notes: This plot shows the output noise spectrum of the reference.
[click] At low frequencies, the plot shows the 1/f noise component of the reference.
[click] At higher frequencies, the flat portion of the curve shows the broadband noise of the reference. Broadband noise can be minimized by filtering the reference. One approach is to simply add a larger filter capacitor at the output of the reference.
This graph shows how broadband noise is minimized using larger output capacitors.
[click] Another approach to minimizing broadband noise is to use an RC filter with a reference buffer. In this case it is possible to set the cutoff frequency very low to minimize the broadband noise. Note that this method is not practical without a reference buffer as the series resistance will limit the dynamic output impedance of the reference.

## Slide 10: Output Impedance
REF60xx
Low output impedance across frequency allows the reference to respond to rapid current transients.
Notes: One specification that is important to driving the SAR ADC reference input is the output impedance. Unfortunately, many references do not provide information on this specification. The graph shown here is for the REF60xx series reference. You can see in this example that the output impedance is very low out to 1MHz. This reference was specifically optimized to drive reference inputs so this output impedance curve is a good bench mark for comparison.

## Slide 11: Important reference specification
Reference Specification | Can Calibrate? | Mitigation
Initial Error | Yes. | Can be calibrated out.
Solder shift | Yes. | Can be calibrated out. Can be minimized by carful PCB design and assembly.
Long Term Shift | Yes. | Periodic calibrations can minimize this error. Also, accelerated aging (i.e. baking the system) can be used to eliminate the largest shifts.
Temperature Drift | No. | Very difficult to calibrate. Choose best temperature drift for your application.
Broadband Noise | -na- | Filtering can minimize broadband noise.
1/f Noise |  | Not practical to eliminate. Choose the lowest 0.1Hz to 10Hz device
Output impedance | -na- | Choose references with low output impedance across frequency for driving dynamic loads like SAR reference inputs.
Notes: This table summarizes the different reference error sources that we discussed. The key point to remember for each error source is how the error can be mitigated. The initial error and solder shift can be calibrated out. Temperature drift on the other hand, is difficult to calibrate out so this is usually minimized by selecting a device with specifications that meet your requirements. Long term shift, or aging, can be eliminated by periotic calibrations. Also, you can bake the product to accelerate aging and eliminate the largest shifts. Broadband noise can be minimized using filtering, but 1/f noise can only be minimized by selecting a device with good specifications. Finally, low wideband output impedance is important for applications with dynamic loads like SAR reference inputs. The only way to “mitigate” a non-optimal output impedance is to use a reference buffer. Otherwise, just select a device with low, wide bandwidth output impedance.

## Slide 12: Thanks for your time! Please try the quiz.
Notes: That concludes this video – thank you for watching! Please try the quiz to check your understanding of this video’s content.

## Slide 13: Overview of Reference Drive Topologies TIPL 4502 TI Precision Labs – ADCs
Created by Luis Chioye, Art Kay
Presented by Cynthia Sosa
Notes: Hello, and welcome to the TI Precision Lab series covering voltage reference considerations for SAR ADCs. The goal of this section is to take a close look at the importance of having a wide bandwidth buffer between the reference and the ADC input. We will compare and contrast measured results showing the importance of the buffer. Finally, we will highlight some devices with integrated buffers that can be used to simplify your design.

## Slide 14
Agenda
Reference Performance Specifications:
  Initial Accuracy, Drift, Long Term Drift, and Noise
Overview of SAR REF Drive Topologies:
Standalone Reference vs. Buffered Reference
  SAR ADCs with Internal Reference Buffer
SAR REF Input Overview: The Capacitive DAC (CDAC)
Build TINA REF Input Model for a SAR:
Discrete Charge Model
TI Device Specific Model
SAR REF Drive Circuit Design:
Reference Bypass Capacitor
Reference Buffer Stability and Compensation
Notes: In this video, an overview of different SAR reference drive circuit topologies will be presented. First we will define what we mean by reference buffer. Then we will compare a buffered and unbuffered reference. Finally, we will also look at how some devices integrate reference buffers.

## Slide 15: ADC Reference input is not a high impedance
REFIN Input
Current Transients (mA)
Convert Start Signal
Initial Current Transient
Bit Transients (12 for 12-Bit Converter)
Pulses are nanoseconds from each other
ADS8028

## Slide 16: What is a reference buffer?
Wide bandwidth
Low output impedance across frequency
Capable of sourcing and relatively large currents (e.g. ±10mA)
Good DC specifications (i.e. offset, and Temperature Drift)
May be integrated in the reference, or an external amplifier
Notes: The term reference buffer is used to describe a wide bandwidth amplifier used between the reference and the ADC input. The reference buffer responds to the very fast transient current requirements on the switched capacitor ADC reference input. Note that multiple transients from the reference input can be milliamps in amplitude and are spaced in nano-seconds. Besides having wide bandwidth, a good reference buffer needs to have low output impedance across frequency. Furthermore, it needs to be capable of sourcing and sinking relatively large currents, for example ±10mA. Also, note that any reference buffer will need to have good dc performance. That is, the reference buffer will need to have low offset and offset drift. Finally, the reference buffer can be integrated into the reference or may be an external amplifier. The figure shown here demonstrates an external amplifier buffer. Note that all series voltage references will have some type of internal buffer. In many cases, however, this buffer doesn’t have sufficient bandwidth or current drive capability to directly drive an ADC reference input and achieve good settling performance. Thus, when we refer to a “reference buffer” in this presentation, we are specifically referring to a wide bandwidth buffer that is capable of driving a SAR ADC reference input and achieving good settling. The figure below shows a reference buffered by an external wide bandwidth amplifier.

## Slide 17: When is the reference buffer required?
Reference buffer requirement depends on:
  Output drive capability of reference.
  SAR reference input current demand:
    Function of sampling rate and resolution performance of ADC.
  Higher resolution devices more challenging
Check device datasheet guidelines for reference drive circuit.
  Detailed analysis and simulation may be required to verify.
Notes: The next logical question you might ask is “when is a reference buffer required?”. The buffer requirement depends primarily on two factors: the output drive capability of the reference, and the SAR ADC input current demand.
The SAR REF input current demand tends to be a function of the sampling rate and resolution of the ADC. As the sampling rate increases, the “average” reference input current demand will tend to increase. It is important to understand that the reference input current in the traditional SAR is not a static or constant DC current. Rather, the reference input current consists of fast current transients as the binary weighted bit decisions are made. We will cover more details on this later.
Also note that higher resolution SARs will tend to have more demanding reference specifications. This is because the reference voltage input must remain stable and settle to less than one LSB. For example, a 12-Bit ADC using a 5V reference has a least significant bit weight (LSB) of 1.22mV. In contrast, a 16-Bit device has a smaller bit weight resolution of 76 microvolts, imposing a much more stringent requirement.
When designing a data acquisition system the first step is to check the SAR ADC datasheet for guidelines; most ADC datasheets provide a circuit recommendation optimal to support the maximum data rate. However, depending on the sampling rate requirements, and the reference device used in the application, a detailed analysis or simulation may be required to verify that you will meet the performance target.
Let’s look at a few example circuits. First, lets consider a case where performance is degraded because we didn’t use a reference buffer.

## Slide 18: Performance limitations from unbuffered reference
REF5050 output doesn’t have sufficient bandwidth to respond to ADC reference input transients
Notes: Here we show the REF5050 connected to the ADS8860 16bit 1Msps converter. In this example we will see degraded performance because the reference doesn’t have sufficient bandwidth to respond to the transient current demand of the ADC. In this circuit, the REF5050 device is not able to recover and recharge the bypass capacitor between conversions, and voltage droop and settling errors occur at the SAR REFP input between conversions. These voltage reference settling errors cause linearity and distortion in the conversion results. Note that the sampling rate and resolution have an impact on the reference bandwidth requirement. In this example the ADS8860 is a 1Msps 16 bit ADC which is relatively fast and high resolution, so a wide bandwidth buffer is needed. Let’s take a look at the measured results.

## Slide 19: Unbuffered Reference: ADS8860 + REF5050
Harmonics / distortion worse than specification
Degraded performance
 | SNR (dB) | THD (dB)
Specification | 93 | -108
REF5050 | 92.3 | -91.3
Notes: This plot shows the FFT results of the ADS8860 sampling at 1-MSPS with the REF5050 reference driving the ADC without a buffer. Comparing the SNR and THD to the specifications you can see that the measured THD of -91.3dB is significantly degraded compared to the expected THD of -108dB. Also, if you look at the FFT you will notice a relatively large second and third harmonics. In the next slide we will resolve this issue by placing a reference buffer with enough bandwidth and low output impedance to drive the reference input.

## Slide 20: Performance improvement using buffered reference
Wide bandwidth integrated buffer handles transients from ADC reference input
Notes: This figure shows a very similar configuration to the last circuit. The data converter and test conditions are the same for both circuits. The only change is that the REF5050 was replaced with the REF6050. Both devices are precision low noise low drift references with very similar specifications. The main difference between the two references is that a wide bandwidth buffer is used at the output of the REF6050. This buffer has low output impedance across frequency and was optimized to drive a switched capacitor reference input. One key benefit of using a device like the REF6050 is that the PCB board area, and complexity is reduced compared to a discreet reference buffer. In fact, both the REF6050 and REF5050 are pin-for-pin compatible, so let’s see how the ADS8860 performance is impacted by changing the reference.

## Slide 21: Buffered vs. Unbuffered.
 | SNR (dB) | THD (dB)
Specification | 93 | -108
REF5050 | 92.3 | -91.3
REF6050 | 92.9 | -107.6
REF5050+ADS8860
1Msps
REF6050+ADS8860
1Msps
Notes: Here we compare the ADS8860 performance with the two different references. First note that the data converter specifications are 93dB for SNR and -108dB for THD. Ideally, our measured system should meet the data sheet specifications. If it doesn’t meet the specification, it is likely that some external component such as the reference is the limiting factor.
The plot on the left shows the performance with the REF5050. Although this reference is a precision low noise low drift reference you can see that the THD is significantly degraded compared to the specification. According to the specification we should get -108dB but we only get 91.3dB. The plot on the right, on the other hand, shows the performance for the REF6050. Comparing the FFTs you can see that the harmonics for this system is significantly improved. Also, comparing the measured SNR and THD to the specification you can see that we are very close to the specification. The main point here is that the reference with the wide bandwidth buffer is able to achieve a better distortion specification because it can settle to the large rapid transients at the reference input. Next let’s look at how sampling rate can impact the reference drive requirements.

## Slide 22: Buffered reference isn’t always required
REF50xx
Sink/source current capability
16-B SAR ADC
fS (max) = 250kSPS
Large Bypass
Capacitors
Notes: Here we show an example where a wide bandwidth reference isn’t always needed. Recall from the previous example that the REF5050 did not have an integrated wide bandwidth output buffer, so we saw degraded performance when this reference was used at 1Msps. Here we are using a different data converter. The ADS8326 is a 16-bit device with a maximum sampling rate of 250ksps. Using this data converter with the REF5050 actually yields very good performance. This is because the relative slow sampling rate of 250kSPS allows the REF5050 (is able) to completely recharge the bypass capacitors within the time interval between conversions. For best performance, is important to use a relatively large bypass capacitor. The REF5050 is capable of driving large capacitive loads. Also, choosing a reference that can source and sink significant current is important. The REF5050 for example can source ±10mA. Be sure to check the capacitive drive and output current specification in the reference data sheet as some references will be limited. Note that for this example configuration, the ads8326 can meet the data sheet THD and SNR specification without a wide band buffer.

## Slide 23: Composite Amplifier Topology
OPA378 chopper op-amp as an input stage for excellent low drift and DC stability of buffer.
High-Bandwidth output buffer (OPA625) provides and wide bandwidth and low-output impedance to drive the SAR REFP input
Simple Buffer
Composite Buffer
Notes: In cases where you need a reference buffer, from a simplicity perspective, you would like to use a simple op amp unity gain follower as shown on the left. The problem with the simple design is that the buffer needs to have both good AC and DC characteristics. The offset and offset drift need to be low or the precision of the reference is lost. Also, the bandwidth needs to be high and the output impedance needs to be low and flat across frequency. Most amplifiers are either optimized for AC or DC characteristics so finding the ideal reference buffer can be challenging. One way to solve this problem is to use a composite amplifier. A composite amplifier will use two op amps. One that is optimized for good DC operation an the other that is optimized for bandwidth. One example of a composite amplifier is shown on the right. The OPA378 input amplifier is a chopper amplifier. Chopper amplifiers offer excellent DC precision as they continuously self calibrate the offset and drift. In this case the offset is 50µV and the drift is 0.25µV/C. However, the bandwidth and output impedance for this device are not sufficient for driving the reference input. The OPA625 output amplifier offers high bandwidth of 125MHz with very low output impedance over frequency, capable of driving the SAR ADC dynamic load. Since the OPA378 is used at the input in the feedback loop, the offset and drift errors of the OPA625 are corrected, and the OPA625 at the output of the composite can still respond to the rapid transients from the ADC. So, the composite essentially gets the best of both worlds. It has the good DC characteristics from the OPA378, and the good AC characteristics from the OPA625. After inspecting this circuit you might dissuaded from using it because of its complexity; however, to achieve the highest performance at high data rates this may be required. Frequently, the ADC data sheet will provide recommendations for ADC drive circuit, so check the data sheet before choosing your topology.

## Slide 24: Device with internal reference buffer: ADS89xxB
Notes: Many modern SAR ADC’s incorporate internal references and reference buffers. This slide shows the high performance ADS8900B, a SAR ADC offering very high 20-B resolution at a fast sampling rate of 1-MSPS. In this case the output of the voltage reference connects to a high impedance buffer input. Thus, the voltage reference doesn’t see any transients from the ADC. This convenient feature allows multiple ADC devices to be connected to a single voltage reference. This approach simplifies calibration and reduces the system complexity. Note that each device requires a local bypass capacitor at the reference buffer output. Another nice benefit of the reference buffer is that the reference can have a tradition RC noise reduction filter connected to it’s output. If you look at this example, the 1k and 10uF filter limit the noise to 159Hz. It is not possible to use this kind of filter without the buffer as the 1kohm resistor would limit its ability to respond to the large current transients from the reference input. In general, modern ADCs tend to integrate more of these advanced features simplifying your design and reducing PCB area.

## Slide 25: Thanks for your time! Please try the quiz.
Notes: That concludes this video – thank you for watching! Please try the quiz to check your understanding of this video’s content.

## Slide 26: SAR REF Input: The Capacitive DAC (CDAC) TIPL 4503 TI Precision Labs – ADCs
Created by Luis Chioye
Presented by Cynthia Sosa
Notes: Hello, and welcome to the TI Precision Lab series covering voltage reference considerations for SAR ADCs. The goal of this section is to give a general overview of the SAR ADC internal circuitry. In particular we will focus on the capacitive DAC, called CDAC. Understanding the CDAC will show us the mechanism that creates the current transients on the reference input of SAR ADCs.

## Slide 27
Agenda
Reference Performance Specifications:
  Initial Accuracy, Drift, Long Term Drift, and Noise
Overview of SAR REF Drive Topologies:
Reference standalone VS Buffered Reference
  SAR ADCs with Internal Reference Buffer
SAR REF Input Overview: The Capacitive DAC (CDAC)
Build TINA REF Input Model for a SAR:
Discrete Charge Model
TI Device Specific Model
SAR REF Drive Circuit Design:
Reference Bypass Capacitor
Reference Buffer Stability and Compensation
Notes: Here we have the agenda for the entire reference series. In this section the internal Capacitive Digital to Analog Converter (or CDAC) and comparator block functionality is covered and an explanation of the SAR reference input behavior is presented. Let’s get started…

## Slide 28: Acquisition phase
ADS8330
RSH
150Ω
Notes: A SAR data converter conversion cycle can be broken into two different phases; the acquisition phase and the conversion phase. During the acquisition phase, the external driving circuit charges the converter’s internal sample and hold capacitor to a voltage equal to the amplifier’s output voltage. Switch S1 is closed at the beginning of the acquisition phase, and switch S1 opens at the end of the acquisition phase when the conversion process starts. The plot at the right shows the sample and hold voltage as it charges to the target voltage during the acquisition period.

## Slide 29: Conversion Phase
CDAC
Reference input
Notes: This slide shows the conversion phase. At the start of the conversion phase the switch S1 opens and the input voltage is stored on the sample and hold capacitor. In this example, you can see that 2.6V is stored on the sample and hold capacitor.
The SAR ADC operates by using a binary search algorithm to converge to the input signal. At the heart of the SAR ADC architecture, the capacitive DAC (or CDAC) generates binary weighted voltages as a function of the reference input, and a high-speed comparator is used to perform the bit decisions.
In this example, we cover a five bit converter, so five comparisons will be done starting with the most significant bit, or MSB, and moving to the least significant bit, or LSB. The SAR ADC converter performs a bit decision every conversion clock cycle, where the CDAC sets a binary weighted voltage at the input of the comparator. The plot at the right shows the binary weighted CDAC voltages as the binary search algorithm converges to the input voltage.
Next, we will show the internal CDAC circuit an provide an overview on how it functions. This overview will give better insight into the origin of the current transients generated at the ADC reference input.

## Slide 30: SAR ADC Architecture: The Capacitive DAC (1)
Based on TLV- Series
SAR ADCs
Notes: Here we show the internal circuit for the CDAC. The CDAC consists of an array of N capacitors with binary weighted values plus one additional LSB capacitor.
This example shows a 5-bit SAR ADC converter. All capacitors have binary weighted values: C, C/2,C/4 to the least significant weight of C/16. The last two capacitors having the value C/16 are connected so that the total parallel capacitance of the capacitor array is 2C.
The conversion process starts with the acquisition phase, where switch SA is closed and switch SB is connected to the input voltage of +2.6V. The input voltage VIN is applied at the ADC input, and a total charge of Q = -2C x VIN is stored in the capacitor array.
The equivalent figure at the right shows the simplification of the input circuit. To reiterate, the capacitance 2C is the parallel combination of all the conversion capacitors. The input voltage is applied directly across the capacitor 2C, and the charge -2C x VIN is stored.

## Slide 31: SAR ADC Architecture: The Capacitive DAC (2)
Based on TLV- Series
SAR ADCs
Notes: After the acquisition phase ends, the conversion phase starts with the hold mode.
During the hold mode, switch SA opens, and all binary weighted capacitor switches S1 to S5’ are connected to GND resulting in a voltage of VC=-2.6V at the comparator input. Notice, the voltage is inverted, since the opposite side of the binary weighted capacitor its now grounded.
Switch SB disconnects from the input pin and connects to the reference voltage input. Again, look at the figure on the right for a simplified view of the circuit.

## Slide 32: SAR ADC Architecture: The Capacitive DAC (3)
Based on TLV- Series
SAR ADCs
Notes: The conversion process is performed by charge redistribution. The conversion sequence starts with the Most Significant Bit (MSB) decision.
During this step, Switch S1 connects the largest capacitor (C) to the reference voltage (VREF). Capacitor C forms a 1-to-1 capacitance divider with the rest of the binary weighted capacitor.
The simplified circuit on the right shows the comparator voltage becomes VC = -VIN + VREF/2 or VC=-0.552V.
Since the voltage at the inverting comparator input is negative, the comparator output goes high, providing the most significant bit decision: MSB=1.
Notice that the reference inputs sees an effective switch capacitance load of C/2 when the MSB decision is made. This capacitive load will cause a rapid transient current spike when it is connected.

## Slide 33: SAR ADC Architecture: The Capacitive DAC (4)
Based on TLV- Series
SAR ADCs
Notes: (4) The conversion process continues with the MSB-1 decision.
The second bit decision step connects C/2 to VREF using switch S2. Since the first bit decision step resulted in MSB=1, switch S1 is turned to GND again to discharge C. Capacitor C/2 forms a 3 to 4 capacitance divider with the rest of the binary weighted capacitor as shown on the simplified circuit on the right. The comparator voltage becomes VC = -VIN + ¾ VREF. The resulting voltage at the inverting comparator input is positive, and the comparator output goes low, providing the second bit decision: MSB-1=0.

## Slide 34: SAR ADC Architecture: The Capacitive DAC (5)
Based on TLV- Series
SAR ADCs
Notes: (5)During the third bit decision, C/4 connects to VREF using switch S3. Since the previous bit decision step was ‘0’, switch S2 is left connected to VREF. The parallel combination of capacitors C/4 and C/2 form a 5 to 8 capacitance divider with the rest of the binary weighted capacitor as shown on the simplified circuit on the right.
The resulting voltage at the inverting comparator input is negative, and the comparator output goes high, resulting on the third bit decision equal to 1.

## Slide 35: SAR ADC Architecture: The Capacitive DAC (6)
Based on TLV- Series
SAR ADCs
Notes: (6)In similar fashion, during the 4th bit decision, C/8 connects to VREF using switch S4. Since the last bit decision step was high, switch S3 is turned to GND again to discharge C/4.
The simplified circuit on the right shows the comparator voltage becomes VC = -VIN + 11/16*VREF. The resulting voltage at the inverting comparator input is positive, and the comparator output goes low, resulting on the fourth bit decision equal to 0.

## Slide 36: SAR ADC Architecture: The Capacitive DAC (7)
Based on TLV- Series
SAR ADCs
Thomas Kulgelstad:
The operation of the SAR-ADC based on charge redistribution,
Analog Applications Journal (slyt176), Texas Instruments, February 2000.
Notes: (7) At the last step, the Least significant bit (LSB) or fifth bit decision is made by switching in C/16 using switch S5 as shown in the circuit. The resulting five bit conversion result is 10100.
During the conversion phase, the reference input sees a dynamic switch capacitance load, as the binary weighted bit decisions are made during each conversion clock cycle.
Note, the reference input voltage must remain stable as the binary weighted capacitor array is charged.
It is important to highlight, the worst case dynamic current transients tend to occur during the most significant bit decisions, when the largest capacitor values are switching.

## Slide 37: SAR ADC Architecture
The Reference is sampled several times during each conversion
High-current transients (~10’s mA range) are present in this REF input where the ADC’s internal capacitor array is switched and charged as the bit decisions are made.
Current transient pulses occur only a few nanoseconds away of each other during conversion.
REFIN Input
Current Transients (mA)
Convert Start Signal
Initial Current Transient
Bit Transients (12 for 12-Bit Converter)
Pulses are nanoseconds from each other
ADS8028
Notes: On the SAR ADC architecture, the reference input is sampled several times during the conversion phase, the reference is sampled once for every binary weighted bit decision during each conversion clock cycle.
This plot shows the reference current transients seen at the reference input of the ADS8028, a 12-Bit ADC.
[click] As the conversion process start, a large initial current transient occurs,
[click] followed by 12 current transients for every bit decision.
[click] Although the “average” current consumption at the Reference input may only be in the range of a few 100s of microAmps, the fast current transients occurring at this pin may be 10’s of mA in magnitude.
The current transients occur every conversion clock cycle as the binary weighted conversion results are made.
[click] In a typical SAR ADC, the conversion clock frequency is in the range of 10s of MegaHertz, and the fast current transient pulses are only a few nanoseconds away from each other.
This is the reason the reference drive circuit requires a relatively large bypass capacitor to restore charge to the binary weighted capacitor array and the reference driver circuit must be capable of driving the dynamic capacitive load while keeping the reference input stable.

## Slide 38: Thanks for your time! Please try the quiz.
Notes: That concludes this video – thank you for watching! Please try the quiz to check your understanding of this video’s content.

## Slide 39
Agenda
Reference Performance Specifications:
  Initial Accuracy, Drift, Long Term Drift, and Noise
Overview of SAR REF Drive Topologies:
Reference standalone VS Buffered Reference
  SAR ADCs with Internal Reference Buffer
SAR REF Input Overview: The Capacitive DAC (CDAC)
Build TINA REF Input Model for a SAR:
Discrete Charge Model
TI Device Specific Model
SAR REF Drive Circuit Design:
Reference Bypass Capacitor
Reference Buffer Stability and Compensation

## Slide 40: Important Datasheet Parameters
Example: ADS8881 18-B, 1-MSPS, Fully-Differential Input ADC
CSH
Sampling Rate = 1-MSPS
Throughput Period: tcycle=1µs
Conversion Time: tconv=710ns (max)
Input Equivalent Circuit: RSH / CSH

## Slide 41: Estimate REFIN Capacitive Load:
Simplify/Redraw as SE circuit
NOTE: REFIN Cmsb not always ¼ Csh; depends on device Architecture!!
This is a first order estimate.

## Slide 42: Conversion Period and Conversion Clock Timing
Estimate internal Conversion Clock period
ADS8881 uses an internal conversion clock (CCLK)
Datasheet Specifies tconv max

## Slide 43: TINA SPICE Equivalent Model
CLOAD = 1/4 CSH= 13.75pF
RSH
Conversion Clock (CCLK)
tcclk =40ns (25MHz)
tCNTL =40ns*8=320ns (mask)
4-8 transients (1*CSH, 2*CSH)
(mask only 8 transients per conv)
Resets
CREF to 0V

## Slide 44: TINA SPICE Equivalent Model

## Slide 45: Example simulation: transient results
Voltage across CREF Load Cap
Transient Current into REFIN Pin
REFIN Pin voltage
REFIN error Voltage
CNTL switch signal
(8xCREF Loads per Conv)
Switch control signal Conv Clock (CCLK)
Switch control signal
RESETs CREF voltage
Notes: Now we are prepared to look at ADC settling using transient analysis. To do this select “Analysis>Transient” in the TINA SPICE menu. Next select a time range that allows you to examine a few conversions. For this example the sampling rate is 1MHz, so we run the transient for 2us for two full cycles. The transient results for this example include the voltage across the charge bucket filter, VCfilt, the voltage on the internal ADC sampling capacitor, VCsh, the error in ADC settling ,Verror, the Op Amp output, Voa, the acquisition control signal, tacq, and the conversion reset signal, tconv.

## Slide 46: Key Result: Error Signal
Check Settling
Look at droop error over many conversions
Droop < ~1LSB

## Slide 47: Key Result: Error Signal
Check Settling
Look at droop error over many conversions
~500 conversions
Droop 7mV >> ~1LSB
(not Optimal)
Example with excessive droop
OPA320 Stable driving load,
but unable to recover at 1-MSPS.

## Slide 48: Key Result: Average Current
Average Current per Datasheet spec is 300µA
Average Current per simulation spec is 530µA
Simulation Current
Meter
Useful to compare sim vs datasheet
Simulator calculates average current and displays value

## Slide 49: Key Result: Average Current
Average Current per Datasheet spec is 300µA
Average Current per simulation spec is 530µA
Simulation Current
Meter
Useful to compare sim vs datasheet
1. Click/select the Current transient curve
2.Select Process…  Averages… on Top Menu.

## Slide 50
Agenda
Reference Performance Specifications:
  Initial Accuracy, Drift, Long Term Drift, Noise and Output Drive
Overview of SAR REF Drive Topologies:
Reference standalone VS Buffered Reference
  SAR ADCs with Internal Reference Buffer
SAR REF Input Overview: The Capacitive DAC (CDAC)
Build TINA REF Input Model for a SAR:
Discrete Charge Model
TI Device Specific Model
SAR REF Drive Circuit Design:
Reference Bypass Capacitor
Reference Buffer Stability and Compensation

## Slide 51: REFIN TI Device Specific Model
  Device specific: model closer to device topology
  Uses variable weighted switching capacitor load
  Behavior may be more accurate and/or closer to real silicon, at the cost of slower simulations/circuit complexity

## Slide 52: REFIN TI Device Specific Model: transient results
Convert Strat signal
Transient Current into REFIN Pin
ADC Input Signal
REF voltage
REFIN error Voltage
Voltage across CSH
Notes: Now we are prepared to look at ADC settling using transient analysis. To do this select “Analysis>Transient” in the TINA SPICE menu. Next select a time range that allows you to examine a few conversions. For this example the sampling rate is 1MHz, so we run the transient for 2us for two full cycles. The transient results for this example include the voltage across the charge bucket filter, VCfilt, the voltage on the internal ADC sampling capacitor, VCsh, the error in ADC settling ,Verror, the Op Amp output, Voa, the acquisition control signal, tacq, and the conversion reset signal, tconv.

## Slide 53: SAR ADC Reference Drive Hands-on TI Precision Labs – ADCs
by Luis Chioye, Dale Li, Art Kay
Presented by Cynthia Sosa
Notes: Hello and welcome to the TI Precision labs hands-on experiment covering different techniques for driving the ADC reference input. In this experiment we compare and contrast how the performance of an ADC is effected by two different references. One reference has an integrated wide bandwidth buffer and the other reference has a lower bandwidth output. We will look at the systems AC and DC performance with the two different references.

## Slide 54: Circuits under test
ADS8860 Data Sheet (1Msps) |  |  |  | 
Parameter | Min | Typ | Max | Unit
SNR | 92 | 93 |  | dB
THD (1kHz) |  | -108 |  | dB
Circuits are identical except for the reference.
Should meet ADS8860 data sheet specifications
Notes: For this experiment we will measure the AC and DC performance of the ADS8860 with two different references. Both references are low noise precision references, but the REF5050 has a low bandwidth output and the REF6050 has a high bandwidth output. Note that the bandwidth refers to the ability of the reference to respond to transients. The reference input of the ADC has significant fast moving current requirements. In this experiment we will see that the system using the wide bandwidth reference has better AC and DC performance because the reference can respond to the ADCs rapid transient current requirements. For AC performance our goal is to get the ADS8860 to meet data sheet SNR and THD. For DC performance we look at settling of the reference after startup. Ideally, the reference is fully settled during the first conversion and the ADC does not have a transient start up error. Notice that Channel 1 is connected to the low bandwidth REF5050 and Channel 2 is connected to the wide bandwidth REF6050. Now let’s look at a summary of the experiments that we will do.

## Slide 55: Expected results
 |  |  | Measured Results |  |  | Example Measurements |  | 
Reference |  | Sampling Rate | SNR (dB) | THD (dB) | Reference settling Vin = 4.9V DC | SNR (dB) | THD (dB) | Reference Droop Vin = 4.9V DC (codes)
 | Data Sheet Typical Specification | 1M | 93 | -108 |  | 93 | -108 | 
1 | REF5050 | 1M |  |  |  | 92.2 | -91.4 | 10.9
2 | REF5050 | 500k |  |  |  | 92.1 | -97.1 | 5.35
3 | REF5050 | 100k |  |  |  | 92.0 | -108.2 | 1.4
4 | REF6050 | 1M |  |  |  | 92.6 | -108.1 | 1.15
5 | REF6050 | 500k |  |  |  | 92.7 | -109.3 | 0.8
6 | REF6050 | 100k |  |  |  | 92.3 | -108.1 | 0.8
Notes: The table summarizes the experiments that we will do throughout this experiment. Your measurements will go under the “measured results” section and they should be comparable to the “example measurements” shown at the right. The experiments 1 to 3 are done with the REF5050, and experiments 4 to 6 are done with the REF6050. The experiments will be done at three different sampling rates, and you will see that performance is limited at higher sampling rates for the ADC using REF5050. To test AC performance we will look at SNR and THD, and to test DC performance we will look at reference settling. Print a copy of this table and fill it out as we progress through the experiments. Let’s get started.

## Slide 56: Connect the hardware
Notes: For this experiment we will first use channel 1 on the precision labs hardware. First, set the jumpers as shown here. Connect the PSI to the PLABS board using the SMA cable, and connect the PHI to the PLABS channel 1 connector. Next, connect the USB cables to the computer. Finally, install OPA320_Goodfilter1 in the socket. Be careful to install this amplifier right side up with the label at the bottom of the coupon card. Pause and connect the hardware.

## Slide 57: Start & Setup the Plabs-SAR-EVM Software
3. The green “HW Connected” and teal “PSI Controls” indicate good hardware communication.
2. Select “SAR EVM”
1. Select “Plabs-Toolkit”
Notes: Now let’s start the software by selecting the “Plabs-SAR-EVM” icon from the “Start > All Programs” menu. Once the software is running you should notice the green “HW connected” message at the bottom of the software. Also if the PSI controls are colored teal the hardware communications is working. Pause and get the EVM software started.

## Slide 58: REF5050: 1Msps, THD, SNR
4. Amplitude = 4.9V
Offset = 2.5V
Frequency = 2kHz
Press power on.
1. Select spectral analysis
3. Press Mark Harmonics
2. Sampling Rate 1Msps
Press “Capture”
Record AC performance
 | fs | SNR (dB) | THD (dB)
Spec. | 1M | 93 | -108
REF5050 | 1M | 92.2 | -91.4
Notes: Now let’s measure the AC performance for the system using the low bandwidth REF5050. First, select “spectral analysis” on the “pages” menu. Next, make sure the sampling rate is 1Msps. Next check the “mark Harmonics” option. This will zoom in on the distortion components once we capture the waveform. Now enter the amplitude, offset, and frequency into the PSI Controls. In this case the amplitude is near full scale at 4.9V, the offset is half full scale at 2.5V, and the frequency is 2kHz. Press the power button to enable the PSI output. Note that it is colored teal when the output is enabled and red when it off. Next press capture and record the AC performance. When using REF5050 at 1Msps you see that THD is significantly degraded compared to the data sheet specifications. In this example the THD is -91.4dB whereas the data sheet specification is -108dB. Also, you can see in the FFT that the harmonics are quite large. Pause and measure the AC performance.

## Slide 59: REF5050: 500ksps, THD, SNR
Press “Capture”
Record AC performance
1. Sampling Rate 500ksps
 | fs | SNR (dB) | THD (dB)
Spec. | 1M | 93 | -108
REF5050 | 1M | 92.2 | -91.4
REF5050 | 500k | 92.1 | -97.1
Notes: Here we will make the same measurement at a sampling rate of 500ksps. First, change the sampling rate. Second press capture, and record the AC performance. No other settings need to change. AC performance is improved significantly by reducing the sampling rate. THD improved from -91.4dB to -97.1dB. Note that it still is not equivalent to the data sheet specification of -108dB. Pause and measure the AC performance.

## Slide 60: REF5050: 100ksps, THD, SNR
 | fs | SNR (dB) | THD (dB)
Spec. | 1M | 93 | -108
REF5050 | 1M | 92.2 | -91.4
REF5050 | 500k | 92.1 | -97.1
REF5050 | 100k | 92.0 | -108.2
Press “Capture”
Record AC performance
1. Sampling Rate 100ksps
Notes: Here we will make the same measurement at a sampling rate of 100ksps. First, change the sampling rate. Second press capture, and record the AC performance. The AC performance here is now equivalents to the data sheet specification of -108dB. Further reductions in sampling rate will have little to no effect on performance. Pause and measure the AC performance.

## Slide 61: Reducing sampling rate improves performance
Notes: In the previous slide you noticed that the data converter AC performance improved as sampling rate decreased. This is because the reference has time to fully settle when the sampling rate is decreases. Looking at this graph, you can see at the top of the graph the current drawn by the ADC reference input. There is a large rapid 50mA current transient for each bit decision. Each group of current transients represents one conversion cycle. In this example, we see three conversion cycles. The bottom graph shows the reference voltage output. You can see corresponding transients in the voltage for each current transients. Between each current transient, the reference needs to replenish the charge on the decoupling capacitor and settle to within half of one LSB. Increasing the conversion time increases the time between each bit decision and between each conversion cycle. This provides more time to charge the reference capacitors and substantially improves the settling on the reference. Good reference settling will improve the data converters AC and DC performance.

## Slide 62: Reference Settling Test – How it works
Notes: Next we are going to see how DC performance is impacted by the bandwidth of the reference. To do this we will alternate between converting and having an idle wait time. For a data converter with a constant DC input the reference voltage will go through a transient behavior when conversions start and eventually settle to a steady state value after many conversions. If the conversions are paused for an idle period, the reference will recover to it’s initial state. Then once the conversions resume the initial transient behavior will repeat itself. Thus if you have a system that “goes to sleep”, when it wakes up and begins converting there will be a transient on the reference that introduces and error for the initial samples. We will see that the magnitude and length of this transient is minimized when a wide bandwidth reference like the REF6050 is used.
From a practical perspective this reference settling is an issue in systems that use a “sleep mode” for power saving. This is because the system measured results will have an error during the initial transient. A similar behavior will also happen in multiplexed systems when transitioning from between channels that are at significantly different voltage levels. In this case the different conversion results from the two different channels will cause a significant shift in the ADC reference load current requirement. The shift in load applied to the reference will cause a reference voltage transient similar to the “sleep mode” case.
In this figure you can see the reference voltage and conversion result. Notice that the conversion result will be the inversion of the reference voltage because the conversion result is calculated by dividing by the reference voltage over 2 to the “n” power. We will measure reference settling by just looking at the conversion result.

## Slide 63: Hardware change
Notes: For this experiment we will make a few minor changes to the hardware. First disconnect the SMA cable from the PSI. Next change the jumper to DC_IN and make sure that the voltage jumper is set to 4.9V. So for this experiment, the AC signal is disconnected and we connect the on-board 4.9V DC source. Pause and change the hardware.

## Slide 64: REF5050: 1Msps, Settling
 | fs | Droop (LSB)
REF5050 | 1M | 10.9
1. Sampling Rate 1Msps
Number of sets to average to 20.
4. Start test
Min Interset delay (ms) to 100ms.
5. Reference droop
Reference Settling Analysis
Notes: Now we can measure reference settling for the REF5050 at 1Msps. First set the sampling rate to 1Msps. Next, set the number of averages to 20. This will take 20 measurements and average them to minimize the noise. Third set the “minimum interset delay” to 100ms. This is just the idle time between each set of 10,000 samples. Fourth, press “start test”. This will take a little time as we are averaging 20 readings. Once this is complete you will see the “reference droop” in LSB. This is the magnitude of the transient error on the reference. In this case the error is 10.9 LSB. You can see most of this happens in the first 500 samples. Pause and measure reference settling.

## Slide 65: Zoom in on Settling
Notes: Right click and drag to zoom in on the transient. Now you can see more clearly that this is a normal second order kind of overshoot and ringing. From this picture it is clear that you will not get the best accuracy until after 300 samples. Pause and zoom in on the settling.

## Slide 66: REF5050: 500ksps, Settling
 | fs | Droop (LSB)
REF5050 | 1M | 10.9
REF5050 | 500k | 5.35
1. Sampling Rate 500ksps
2. Start test
3. Reference droop
Notes: Now let’s look at reference settling with the sampling rate set to 500ksps. First set sampling to 500ksps. Next, press start test. After this test completes you will see that the reference droop is significantly better for the 500ksps case. The droop went from 10.9 LSB to 5.35 LSB. Ideally, we would like this error to be less than 1 LSB. So, this is an improvement but still not ideal. Pause and measure settling at 500ksps.

## Slide 67: REF5050: 100ksps, Settling
 | fs | Droop (LSB)
REF5050 | 1M | 10.9
REF5050 | 500k | 5.35
REF5050 | 100k | 1.4
1. Sampling Rate 100ksps
2. Start test
3. Reference droop
Notes: Now let’s check settling at 100ksps. First set the sampling rate to 100ksps. Second press “start test”. After the test completes you will see that the reference droop is about 1.4 LSB. Although this isn't perfect, it is pretty close to the goal of less than 1 LSB. Thus, here we can see that the dc performance is also improved by reducing the sampling rate. This reference settling is especially important when a multiplexer is used as each time the input steps the reference will have to settle again. We will perform the same set of tests with the REF6050, and you will see that the droop for this reference is good at the entire range of sampling rates. Pause and measure settling at 100ksps.

## Slide 68: Disconnect USB and Close Software
Notes: Now we are going to switch to another channel. To do this we need to disconnect power to both the PSI and PHI by disconnecting the USB cable. Also, close the software should be closed by clicking the “x” in the corner.

## Slide 69: Connect the hardware
Notes: Now we will switch to channel 2 on the precision labs hardware. First, set the jumpers as shown here. Note that we are changing back to an AC input with the jumper settings. Connect the PSI to the PLABS board using the SMA cable, and connect the PHI to the PLABS channel 2 connector. Next, connect the USB cables to the computer. Finally, install OPA320_Goodfilter1 in the socket. Be careful to install this amplifier right side up with the label at the bottom of the coupon card. Pause and connect the hardware.

## Slide 70: Start & Setup the Plabs-SAR-EVM Software
2. The green “HW Connected” and teal “PSI Controls” indicate good hardware communication.
1. Select “Plabs-SAR-EVM” from “start>All Programs”
Notes: Now we can re-start the software by selecting the “Plabs-SAR-EVM” icon from the “Start > All Programs” menu. Once the software is running you should notice the green “HW connected” message at the bottom of the software. Also if the PSI controls are colored teal the hardware communications is working. Pause and get the EVM software started.

## Slide 71: REF6050: 1Msps, 500ksps, 100ksps: SNR, THD
4. Amplitude = 4.9V
Offset = 2.5V
Frequency = 2kHz
1. Select spectral analysis
3. Press Mark Harmonics
2. Sampling Rate 1Msps, 500ksps, 100ksps
Press “Capture”
Record AC performance
 | fs | SNR (dB) | THD (dB)
Spec. | 1M | 93 | -108
REF6050 | 1M | 92.6 | -108.1
REF6050 | 500k | 92.7 | -109.3
REF6050 | 100k | 92.3 | -108.1
Notes: Now we will repeat the SNR and THD measurements for the REF6050. Remember the REF6050 has a wide bandwidth integrated buffer so we expect the system using this device to have good performance up to the full 1Msps sampling rate. To perform this test first select “Spectral Analysis” under “Pages”. Next select the sampling rate. For this test we will use 1Msps, 500ksps, and 100ksps sampling rates. Let’s start with 1Msps. Next check the “Mark Harmonics” check box. This will zoom in on the distortion components in the FFT. Next configure the PSI. Set the amplitude to 4.9V, the offset to 2.5V, and the frequency to 2kHz. Finally, press capture and record the AC performance. Pause and measure SNR and THD at all three sampling rates.
The table to the right compares the SNR and THD specification to the measured performance at the different sampling rates. As expected, the measured performance is very close to the specified performance for all the different sampling rates tested. This is because the wide bandwidth reference in the REF6050 settles quickly allowing for low distortion signal capture. Next we will look at the DC performance with the REF6050.

## Slide 72: Hardware change
Notes: For this experiment we will make a few minor changes to the hardware. First disconnect the SMA cable from the PSI. Next change the jumper to DC_IN and make sure that the voltage jumper is set to 4.9V. So for this experiment, the AC signal is disconnected and we connect the on-board 4.9V DC source. Pause and change the hardware.

## Slide 73: REF6050: 1Msps, 500ksps, and 100ksps settling
Sampling Rate 1Msps, 500ksps, and 100ksps
3. Number of sets to average to 20.
5. Start test
4. Min Interset delay (ms) to 100
6. Reference droop
 | fs | Droop (LSB)
REF6050 | 1M | 1.15
REF6050 | 500k | 0.8
REF6050 | 100k | 0.8
Select “reference settling”
Notes: Here we will check the reference settling for the REF6050 by applying a DC 4.9V input and measuring the conversions. When we performed the same experiment using the REF5050 the performance was degraded at higher sampling rates. For this experiment first select “Reference Settling” under “Pages”. Next, select the sampling rate. Here we will run the experiment for 1Msps, 500ksps, and 100ksps. Let’s start with 1Msps. Next enter the number of data sets that we will average to 20. Next set the “Interset” delay to 100ms. This is the delay between each set of points being measured. Finally, press “Start test” and measure the reference droop. The figure shown shows the results for 1Msps. Repeat the experiment for 500ksps, and 100ksps. Pause and measure droop at all three sampling rates.
The table to the right shows the settling performance for the REF5050 system. Ideally the droop should be less than 1 LSB. As expected, the measured performance is very close to this target droop error. This is because the wide bandwidth reference in the REF6050 responds quickly to the transient current requirements and holds the reference voltage constant. Now let’s compare all the measured results.

## Slide 74: Expected results
Your results should show the same trend as the expected result but the specific values will differ.
 |  |  | Measured Results |  |  | Example Measurements |  | 
Reference |  | Sampling Rate | SNR (dB) | THD (dB) | Reference settling Vin = 4.9V DC | SNR (dB) | THD (dB) | Reference Droop Vin = 4.9V DC (codes)
 | Data Sheet Typical Specification | 1M | 93 | -108 |  | 93 | -108 | 
1 | REF5050 | 1M |  |  |  | 92.2 | -91.4 | 10.9
2 | REF5050 | 500k |  |  |  | 92.1 | -97.1 | 5.35
3 | REF5050 | 100k |  |  |  | 92.0 | -108.2 | 1.4
4 | REF6050 | 1M |  |  |  | 92.6 | -108.1 | 1.15
5 | REF6050 | 500k |  |  |  | 92.7 | -109.3 | 0.8
6 | REF6050 | 100k |  |  |  | 92.3 | -108.1 | 0.8
Notes: The table summarizes the measured results for all tests in this experiment. Your measured results should be close to the example measurements. Again note that both the AC and DC performance for the REF6050 is good for all the sampling rates whereas the performance for the REF5050 is only good at lower sampling rates.

## Slide 75
Thanks for your time!
Notes: That concludes the hands-on experiment. I hope this was useful to you. Thanks for your time.